NATIONAL RESEARCH DEVELOPMENT CORPORATION (NRDC)  •  WORK PLAN & MILESTONES

ANNEXURE: QUARTERLY WORK PLAN & MILESTONES

Detailed 6-Month Timeline, Work Packages, Milestone Deliverables and Responsibilities Aligned with ₹1,70,000 Budget

# 1. Project Executive Schedule & Work Breakdown Structure

The proposed Category 2 testing and trial program is structured over a tightly managed 6-month duration (2 Quarters: Month 1 to Month 6). It focuses strictly on quantifiable bench testing, multi-rater concordance evaluation, repeatability benchmarking, and failure-mode analysis.

| Quarter / Month | Work Package (WP) | Detailed Experimental Activities | Responsible Lead |
| --- | --- | --- | --- |
| Q1: Month 1 | WP1: Experimental Jig & Bench Protocol Setup | Procurement of multi-layer silicone pads and calibration targets. Fabrication of rigid overhead camera mount. Laser-micrometer baseline calibration (0.124 mm/px scale validation). | Armaan Samir Jena (Applicant 2 / Engineering Lead) |
| Q1: Month 2 | WP2: Optical Calibration & Lighting Stress Trials | Execution of 60 standardized suture bench trials under 4 controlled lighting levels (300lx, 600lx, 900lx, 1200lx). Quantitative measurement of needle tip localization error and angle variance. | Armaan Samir Jena (Applicant 2) & Research Staff |
| Q1: Month 3 | WP3: Needle Curvature Repeatability Benchmark | Testing repeatability across 4 standard surgical needle geometries (3/8 circle, 1/2 circle, reverse cutting, round-bodied). Mid-term technical review and milestone M1 report compilation. | Dr. VIJAYANANDA (Applicant 1 / Clinical Lead) |
| Q2: Month 4 | WP4: Multi-Rater Faculty Inter-Rater Reliability Trials | Conducting 60 resident simulation trials evaluated concurrently by 3 independent blinded attending surgeons using the Faculty Review Console. Calculation of Intraclass Correlation (ICC) and Bland-Altman limits. | Dr. VIJAYANANDA (Applicant 1) & Faculty Reviewers |
| Q2: Month 5 | WP5: Occlusion & Failure-Mode Stress Analysis | Evaluating tracking drop-out under rapid suture knot cinching, instrument hand-crossing, and fluid visual distortion. Testing automated fallback and recovery latency in cvEngine.ts. | Armaan Samir Jena (Applicant 2) & Simulation Staff |
| Q2: Month 6 | WP6: Statistical Synthesis, Audit & Final Reporting | Independent statistical verification of concordance data. Compilation of final technical dossier, CA audited Utilization Certificate (UC), and institutional commercialization roadmap. | Dr. VIJAYANANDA (Applicant 1) & Armaan Samir Jena (Applicant 2) |

# 2. Milestone Deliverables Matrix

| Milestone | Target Deadline | Verifiable Milestone Output / Deliverable | Funding Linkage |
| --- | --- | --- | --- |
| Milestone M1 | End of Month 2 | Calibrated Bench Setup & Optical Scale Uncertainty Report (0.124 mm/px validation data) | Tranche 1 (₹85,000) |
| Milestone M2 | End of Month 3 | Mid-Term Progress Report & Needle Geometry Repeatability Dataset (60 trials) | Tranche 1 Audit |
| Milestone M3 | End of Month 4 | Multi-Rater Faculty Agreement Dossier (ICC & Bland-Altman Concordance Analysis across 60 resident trials) | Tranche 2 (₹85,000) |
| Milestone M4 | End of Month 5 | Failure-Mode Resilience & Recovery Latency Technical Report | Tranche 2 Review |
| Milestone M5 | End of Month 6 | Final Comprehensive Technical Evaluation Dossier, Audited UC, and Commercialization Plan | Project Completion |

✦ WORK PLAN COMPLIANCE STATEMENT: Every activity outlined in this work plan is directly linked to the ₹1,70,000 itemised budget. No capital expenditure or unbudgeted activities are included. All trials are non-clinical and conducted in simulation laboratories.